## Specification
Specification table for electromagnetic flow meter FLONET FF50xx |  |  |  | Recommended value | Error message / note
Click each white field in column D to turn it off (most of the predefined values ). |  |  |  | The table serves as an attachment to the inquiry or order and for production sector.
Pol. | TYPE IDENTIFICATION |  | Code ON
5 | Connection of the sensor
6 | Flowmeter version and equipment
- | TECHNICAL PARAMETERS | - | -
8 | Dimension of sensor DN/max flow rate
9 | Sensor material
10 | Sensor lining
11 | Sensor protection
12 | Electrodes material
13 | Grounding electrodes
14 | Nominal (maximum allowable) pressure
15 | Max operating medium temperat. (TSmax)
16 | Power supply
17 | Transmitter protection class
18 | Type of measured liquid |  |  |  |  | The above parameters defining the design of the product.
19 | Description of the measured liquid
20 | Grounding rings
21 | Length of cable
- | FLOW METER SETTINGS |  | -
23 | Pulse number
24 | Units of pulse number
25 | Units of volume
26 | Flow units
27 | Number of samples for averaging
28 | Measurement in sensitivity
29 | Displayed languague
30 | Data on display
31 | Flow 100% (l/s)
32 | Current output Flow rate
33 | Output OUT 1
34 | Output OUT 1 functions
35 | Output OUT 2
36 | Output OUT 2 functions
- | COMMUNICATION SETTINGS | - | -
38 | Interface
39 | Address - enter from 1 to 247
40 | Transmission speed
41 | HART® | NOT requested
42 | FLOSET 4.0 with accessories
- | CALIBRATION | - | -
44 | Calibration requirement
- | PURCHASE CONDITIONS | - | -
46 | Number of pieces
47 | Packaging
48 | Delivery
49 | Warranty
- | RELATED REGULATIONS | - | -
51 | Number of manual flowmeter | Es90736K | 1
52 | IDO of flowmeter (given by ELIS)
| Ordering number |  |  | In case of possible copying of the Order No. listed on the left
| The complete order number is not generated until all parameters have been entered |  |  | use function "Insert values"!
#REF!
| Here specify all parameters marked as NON-STANDARD (code "x") |  |  | These parameters (item number - name):
Date
Company
Your inquiry / order number
Item number in the inquiry / order
Customer´s responsible person
ELIS responsible person
Es90163F
Note on D4: FLONET FH30xx is a volume flow meter for conductive liquids in pipes. It enables bi-directional flow measurement with high accuracy over the range of flow velocities (0.025 to 10 m/s). The minimum required conductivity of the measured medium is 10 μS/cm, for demineralized water 20 μS/cm.
Note on C14: Nominal pressure (PN) = maximum allowable pressure (PSmax)
Note on D14: Standard combinations of nominal pressures and dimensions: PN 25 ~ DN 6 to 10 (1/8" to 1/4") PN 40 ~ DN 15 to 50 (1/2"to 2") PN 16 ~ DN 65 to 200 (2 1/2"to 8") PN 10 ~ DN 250 to 700 (10" to 28") PN 6 ~ DN 800 to 1200 (32"to 48")
Note on D19: This data has no predefined values. Provide the most accurate description possible
Note on D21: Do not fill in the compact version!
Note on D23: See sheet "Pulse numbers"
Note on C25: Displayed volume units
Note on C26: Displayed flow units
Note on D33: Q = Q₃
Note on D35: Q = Q₃
Note on D39: 4 = standard
Note on D40: Baud (Bd) is a unit of modulation rate indicating the number of changes in the state of the transmission medium per 1 second.
Note on C42: Accessories: USB / RS-485 converter, connecting cables, User Manual.
Note on D42: Through the FLOSET program, the flow meter can be connected to a computer (PC, laptop, tablet) on which Windows 7 and higher (or Linux, iOS) with JAVA 8u40 and higher is installed.
Note on C43: The calibration and the official verification is icluded.
Note on D44: You will always receive a calibration report when calibrating.
Note on D52: Product IDO (Bill of Material No. incl. execution) is entered here by an authorized ELIS employee, if the product is in the Helios database.
Note on C59: To type in this field, double-click in the field (double click - the cursor appears ...). (New Line = Alt + Enter)

## Data (hidden)
Connection of the sensor | Code |  |  |  | Sloupec pro zápis návrhu úprav
DIN 11851 food connection | 1
Clamp DIN 32676 | 2
Tri Clamp® | 3
Clamp ITE Intertechnik | 4
Non-standard | x
Flowmeter version and equipment |  | Přepočet m³ na US Gal
COMFORT compact | 4 | US gal/m³: | 264.1721
COMFORT remote | 5
Dimension of sensor |  | m³/h | US gal/h | US GPM (US gal/min)
DN10 2.8 m³/h (NPS 3/8" 12.3 US GPM) | 03 | 2.8 | 739.6817 | 12.328
DN15 7.87 m³/h (NPS 1/2" 34.6 US GPM) | 04 | 7.87 | 2079.0341 | 34.6506
DN20 12 m³/h (NPS 3/4" 52.8 US GPM | 05 | 12 | 3170.0646 | 52.8344
DN25 20 m³/h (NPS 1" 88.1 US GPM) | 06 | 20 | 5283.441 | 88.0574
DN32 31.25m³/h (NPS 1¼" 137.6 US GPM) | 07 | 31.25 | 8255.3766 | 137.5896
DN40 50 m³/h (NPS 1½" 220.1 US GPM) | 08 | 50 | 13208.6026 | 220.1434
DN50 79 m³/h (NPS 2" 347.8 US GPM) | 09 | 79 | 20869.5921 | 347.8265
DN65 125 m³/h (NPS 2½" 550.4 US GPM) | 10 | 125 | 33021.5065 | 550.3584
DN80 200 m³/h (NPS 3" 880.6 US GPM) | 11 | 200 | 52834.4105 | 880.5735
DN100 312.5 m³/h (NPS 4" 1 376 US GPM) | 12 | 312.5 | 82553.7664 | 1375.8961
Non-standard | xx
Sensor material
Complete stainless steel 1.4301 | 1
Complete stainless steel 1.4571 | 2
Sensor lining
Teflon PTFE (-20 to +150 °C / -4 to +302 °F) | 4
PVDF (-20 to +120 °C / -4 to 248 °F) | 6
Sensor protection
Standard IP 67 | 1
Non-standard IP 68 | 2
Electrodes material
Stainless steel 1.4571 | 1
Hastelloy C-276 | 2
Platina-Rhodium PtRh20 | 3
Tantalum | 4
Titanium Gr.2 ASTM B348-03 | 5
Non-standard | x
Grounding electrode
YES | 1 | Od DN25
No | 0
Nominal and maximum allowable pressure
PN 10 | 1
Non-standard | x
Maximum operation temperature of medium (TSmax)
50 °C / 122 °F | 1
60 °C / 140 °F | 2
70 °C / 158 °F | 3
80 °C / 176 °F | 4
90 °C / 194 °F | 5
110 °C / 230 °F | 6
120 °C / 248 °F | 7
130 °C / 266 °F | 8
150 °C / 302 °F | 9
Non-standard | x
Power supply
95 to 250 V AC, 45 to 65 Hz | 1
24 V DC ±20 % (19,2 to 28,8 V DC) | 2
Protection of the transmitter
Standard IP 67 | 1
Above standard IP 68 | 2
Measured fluid
Water | 1
Non-standard | x
Description of measured liquid
Není nabízen ze seznamu
Grounding rings
Without grounding rings | 0
Stainless steel 1.4571 | 1
Length of cables
Without cables | 0
5 m (16 ft) (standardly) | 1
10 m (33 ft) | 2
15 m (50 ft) | 3
20 m (66 ft) | 4
30 m (98 ft) | 5
40 m (131 ft) | 6
50 m (164 ft) | 7
100 m (328 ft) | 8
Non-standard | x
Pulse number
NOT requested | 0
Standard (see sheet "Pulse numbers") | 1
Non-standard | x
Units of pulse number
Žádná jednotka (impulzní číslo není požadováno) | 00
l/pulse (standardly) | 01
hl/pulse | 02
milion l/pulse | 03
m³/pulse | 04
US liquid gal/pulse | 05
milion US liquid gal/pulse | 06
Imp gal/pulse | 07
bbl/pulse | 08
bblLiq/pulse | 09
bush US/pulse | 10
in³/pulse | 11
ft³/pulse | 12
yd³/pulse | 13
Non-standard | xx
Units of volume
l | 01
hl | 02
milion l | 03
m³ (standardly) | 04
US liquid gal | 05
milion US liquid gal | 06
Imp gal | 07
bbl | 08
bblLiq | 09
bush US | 10
in³ | 11
ft³ | 12
yd³ | 13
Non-standard units | xx
Units of flow rate
l/s | 01
l/min | 02
l/h | 03
milion l/day | 04
m³/s | 05
m³/min | 06
m³/h (standardly) | 07
m³/day | 08
ft³/s | 09
ft³/min | 10
ft³/h | 11
ft³/day | 12
US liquid gal/s (GPS) | 13
US liquid gal/min (GPM) | 14
US liquid gal/hr (GPH) | 15
US liquid gal/day (GPD) | 16
milion US liquid gal/day | 17
Imp gal/s | 18
Imp gal/min | 19
Imp gal/h | 20
Imp gal/day | 21
bbl/s | 22
bbl/min | 23
bbl/h | 24
bbl/den | 25
Non-standard | xx
Numbers samples for avaraging
25 | 1
50 | 2
100 (standardly) | 3
150 | 4
200 | 5
250 | 6
Non-standard | x
Insensitivity of measurement
50 % Q₁ (standardly) | 1
35 % Q₁ | 2
Non-standard | x
Displayed language
Czech | 01
English | 02
Dutch | 03
Español | 04
Polski | 05
Français | 06
Italiano | 07
Deutsch | 08
По Русски | 09
Português | 10
Non-standard | xx
Display
Total volume (standardly) | 1
Time/date | 2
Operation time | 3
Excitation current | 4
Percentage flow rate | 5
Last error | 6
Flow rate 100% (l/s)
Standard flow Q₃ | 1
Non-standard | x
Current output
NOT requested | 00
Q- to Q+ (Q = Q₃) (standardly) | 01
Q- to Q+ acc to a request | x2
0 to ¦Q¦ (Q = Q₃) | 03
0 ÷ ¦Q¦ acc to a request | x4
Output OUT1
NOT requested | 00
Frequency for Q+ 0 to 10 kHz ~ 0 to Q₃ | 01
Frequency for Q- 0 to 10 kHz ~ 0 to Q₃ | 02
Frequency for ¦Q¦ 0 to 10 kHz ~ 0 to Q₃ | 03
Frequency for Q+ 0 to 1 kHz ~ 0 to Q₃ | 04
Frequency for Q- 0 to 1 kHz ~ 0 to Q₃ | 05
Frequency for ¦Q¦ (standardly) 0 to 1 kHz ~ 0 to Q₃ | 06
Frequency - non-standard (Q ≠ Q₃) | x1
Pulse for Q+ standard pulse width and pulse number, see sheet "Pulse numbers" | 08
Pulse for Q- standard pulse width and pulse number, see sheet "Pulse numbers" | 09
Pulse for ¦Q¦ standard pulse width and pulse number, see sheet "Pulse numbers" | 10
Pulse - non-standard pulses or pulse number | x2
Indication of limit flow Q > Qlimit (Qlimit = Q₃ - closes, if Q > Qlimit) | 12
Indication of limit flow Q¦ > Qlimit (Qlimit = Q₃ - closes, if¦Q¦ > Qlimit) | 13
Indication of stay withour defect (It is closed, if the flow meter is without any defect) | 14
Permanently closed | 15
Function OUT1
Výstup OUT1 není požadován či není pulzní | 0
Positive (standardly) | 1
Negative | 2
Output OUT2
NOT requested | 00
Frequency for Q+ 0 to 10 kHz ~ 0 to Q₃ | 01
Frequency for Q- 0 to 10 kHz ~ 0 to Q₃ | 02
Frequency for ¦Q¦ 0 to 10 kHz ~ 0 to Q₃ | 03
Frequency for Q+ 0 to 1 kHz ~ 0 to Q₃ | 04
Frequency for Q- 0 to 1 kHz ~ 0 to Q₃ | 05
Frequency for ¦Q¦ (standardly) 0 to 1 kHz ~ 0 to Q₃ | 06
Frequency - non-standard (Q ≠ Q₃) | x1
Pulse for Q+ standard pulse width and pulse number, see sheet "Pulse numbers" | 08
Pulse for Q- standard pulse width and pulse number, see sheet "Pulse numbers" | 09
Pulse for ¦Q¦ standard pulse width and pulse number, see sheet "Pulse numbers" | 10
Pulse - non-standard pulses or pulse number | x2
Indication of limit flow > Qlimit (Qlimit = Q₃ - closes, if is Q > Qlimit) | 12
Indication of limit flow Q¦ > Qlimit (Qlimit = Q₃ - closes, if is ¦Q¦ > Qlimit) | 13
Indication of the stay without defect (it is closed, if the flow meter is without any defect) | 14
Permanently closed | 15
Function of OUT2
Výstup OUT2 není požadován či není pulzní | 0
Positive (standardly) | 1
Negative | 2
Interface
Not required | 0
RS-485 MODBUS RTU | 1
Address
Není nabízena ze seznamu
Speed of transfer
300 Bd | 1
600 Bd | 2
1 200 Bd | 3
2 400 Bd | 4
4 800 Bd | 5
9 600 Bd (standardly) | 6
19 200 Bd | 7
38 400 Bd | 8
56 800 Bd | 9
HART® (na listu Specifikace zatím skrytý řádek)
NOT requested | 0
Yes (do not use, still in preparation) | 1
FLOSET 4.0 incl. accessories
NOT requested | 0
Yes | 1
Metrological requests
Without metrological verification | 0
Standard calibration acc to EN ISO 4064-1, accuracy class 2 | 1
Above standard calibration with accuracy ±0,5 % | 2
Above standard calibration with accuracy ±0,2 % | 3
Non - standard metrological request | x
Number of pieces
Není nabízen ze seznamu
Packing
Unpacking | 0
Standard | 1
Non-standard | x
Method of delivery og goods
Personally | 1
By shipping agent on supplier's costs | 2
By shipping agent on buyer's costs | 3
Non-standard | x
Warranty
6 months | 1
12 months (standardly) | 2
18 months | 3
24 months | 4
36 months | 5
Non-standard | x
Note on A70: Nenabízí se, dosazuje se automaticky.
Note on A85: Nenabízí se, dosazuje se automaticky.
Note on A150: Měření vždy OBOUSMĚRNÉ
Note on A195: Indikace místo Diagnostika. Opravit i u ostatních TAB?
Note on A200: Nenabízí se, dosazuje se automaticky.
Note on A221: Nenabízí se, dosazuje se automaticky.
Note on B248: Pro nastavený nulový průtok a referenční podmínky: Relativní chyba Rozsah průtoků ±0,5 % z měřené hodnoty 5–100% Q₄
Note on B249: Pro nastavený nulový průtok a referenční podmínky: Relativní chyba Rozsah průtoků ±0,2 % z měřené hodnoty 10–100% Q₄

## Pulse numbers
Overview of pulse numbers
depending on frequency and dimension
DN [mm] | NPS | Q4 |  | Pulse number [l/pulse]
|  | (v = 10 m/s) |  | for fmax = 100 Hz |  |  | for fmax = 50 Hz (standard) |  |  | for fmax = 10 Hz |  |  | for fmax = 5 Hz
|  | [m3/h] | [l/s] | pulse | f pulse | pulse | pulse | f pulse | pulse | pulse | f pulse | pulse | pulse | f pulse | pulse
|  |  |  | volume | for Q4 | width | volume | for Q4 | width | volume | for Q4 | width | volume | for Q4 | width
|  |  |  | V | f | Pmax | V | f | Pmax | V | f | Pmax | V | f | Pmax
|  |  |  | [l/pulse] | [Hz] | [ms] | [l/pulse] | [Hz] | [ms] | [l/pulse] | [Hz] | [ms] | [l/pulse] | [Hz] | [ms]
10 | 3/8“ | 2.8 | 0.77778 | 0.01 | 77.7778 | 6 | 0.1 | 7.7778 | 64 | 0.1 | 7.7778 | 64 | 1 | 0.77778 | 643
15 | ½“ | 7.87 | 2.1861 | 0.1 | 21.8611 | 23 | 0.1 | 21.8611 | 23 | 1 | 2.1861 | 229 | 1 | 2.1861 | 229
20 | ¾“ | 12 | 3.3333 | 0.1 | 33.3333 | 15 | 0.1 | 33.3333 | 15 | 1 | 3.3333 | 150 | 1 | 3.3333 | 150
25 | 1“ | 20 | 5.5556 | 0.1 | 55.5556 | 9 | 0.1 | 55.5556 | 9 | 1 | 5.5556 | 90 | 1 | 5.5556 | 90
32 | 1 ¼“ | 31.25 | 8.6806 | 0.1 | 86.8056 | 6 | 1 | 8.6806 | 58 | 1 | 8.6806 | 58 | 10 | 0.86806 | 576
40 | 1 ½“ | 50 | 13.8889 | 1 | 13.8889 | 36 | 1 | 13.8889 | 36 | 10 | 1.3889 | 360 | 10 | 1.3889 | 360
50 | 2“ | 79 | 21.9444 | 1 | 21.9444 | 23 | 1 | 21.9444 | 23 | 10 | 2.1944 | 228 | 10 | 2.1944 | 228
65 | 2 ½“ | 125 | 34.7222 | 1 | 34.7222 | 14 | 1 | 34.7222 | 14 | 10 | 3.4722 | 144 | 10 | 3.4722 | 144
80 | 3“ | 200 | 55.5556 | 1 | 55.5556 | 9 | 1 | 55.5556 | 9 | 10 | 5.5556 | 90 | 10 | 5.5556 | 90
100 | 4“ | 312.5 | 86.8056 | 1 | 86.8056 | 6 | 10 | 8.6806 | 58 | 10 | 8.6806 | 58 | 100 | 0.86806 | 576

## Tech (hidden)
Objednací číslo |  |  |  | Počet nezadaných hodnot: | 40 | Nestandardní parametry - souhrn:
|  |  |  | HW část: | 12
Pozice | Kód | Název | Hodnota
0 | FF50 |  |  |  |  | Nestandardní parametry
5 |  | Connection of the sensor | 0
6 |  | Flowmeter version and equipment | 0
- | - | TECHNICAL PARAMETERS | -
8 |  | Dimension of sensor DN/max flow rate | 0
9 |  | Sensor material | 0
12 |  | Electrodes material | 0
10 |  | Sensor lining | 0
11 |  | Sensor protection | 0
13 |  | Grounding electrodes | 0
14 |  | Nominal (maximum allowable) pressure | 0
15 |  | Max operating medium temperat. (TSmax) | 0
16 |  | Power supply | 0
17 |  | Transmitter protection class | 0
18 |  | Type of measured liquid | 0
19 |  | Description of the measured liquid | 0
20 |  | Grounding rings | 0
21 |  | Length of cable | 0
- | - | FLOW METER SETTINGS | 0
23 |  | Pulse number | 0
24 |  | Units of pulse number | 0
25 |  | Units of volume | 0
26 |  | Flow units | 0
27 |  | Number of samples for averaging | 0
28 |  | Measurement in sensitivity | 0
29 |  | Displayed languague | 0
30 |  | Data on display | 0
31 |  | Flow 100% (l/s) | 0
32 |  | Current output Flow rate | 0
33 |  | Output OUT 1 | 0
34 |  | Output OUT 1 functions | 0
35 |  | Output OUT 2 | 0
36 |  | Output OUT 2 functions | 0
- | - | COMMUNICATION SETTINGS | -
38 |  | Interface | 0
39 |  | Address - enter from 1 to 247 | 0
40 |  | Transmission speed | 0
41 |  | HART® | NOT requested
42 |  | FLOSET 4.0 with accessories | 0
- | - | CALIBRATION | -
44 |  | Calibration requirement | 0
- | - | PURCHASE CONDITIONS | -
46 |  | Number of pieces | 0
47 |  | Packaging | 0
48 |  | Delivery | 0
49 |  | Warranty | 0
- | - | RELATED REGULATIONS | -
51 | 1 | Number of manual flowmeter | Es90736K