## Specification
Specification table for electromagnetic flow meter FLONET FN50xx.2 |  |  |  | Recommended value | Error message / note
Click each white field in column D to turn it off (most of the predefined values ). |  |  |  | The table serves as an attachment to the inquiry or order and for production sector.
Pol. | TYPE IDENTIFICATION |  | Code ON
5 | Sensor design
6 | Flowmeter version and equipment
- | TECHNICAL PARAMETERS | - | .2-
8 | Dimension of sensor DN/max flow rate
9 | Sensor flanges
10 | Sensor material, coating
11 | Electrodes material
12 | Sensor lining
13 | Sensor protection
14 | Grounding electrodes
15 | Nominal pressure PN
16 | Max operating medium temperat. (TSmax)
17 | Maximum allowable pressure (PSmax)
18 | Power supply
19 | Transmitter protection class
20 | Type of liquid by PED 2014/68/EU
21 | Category by PED 2014/68/EU
22 | Type of measured liquid |  |  |  |  | The above parameters defining the design of the product.
23 | Description of the measured fluid
24 | Grounding rings
25 | Length of cable
- | FLOW METER SETTINGS |  | -
27 | Pulse number
28 | Units of pulse number
29 | Units of volume
30 | Flow units
31 | Number of samples for averaging
32 | Measurement in sensitivity
33 | Flow 100% (l/s)
34 | Output OUT
- | COMMUNICATION SETTINGS | - | -
36 | Interface
37 | Address - enter from 1 to 247
38 | Transmission speed
39 | FLOSET 4.0 with accessories
- | CALIBRATION | - | -
41 | Calibration
- | PURCHASE CONDITIONS | - | -
43 | Number of pieces
44 | Packaging
45 | Delivery
46 | Warranty
- | RELATED REGULATIONS | - | -
48 | Number of manual flowmeter
49 | IDO of flowmeter (given by ELIS)
| Ordering number |  |  | In case of possible copying of the Order No. listed on the left
| The complete order number is not generated until all parameters have been entered |  |  | use function "Insert values"!
FN50.2------
| Here specify all parameters marked as NON-STANDARD (code "X") |  |  | These parameters (item number - name):
Date
Company
Your inquiry / order number
Item number in the inquiry / order
Customer´s responsible person
ELIS responsible person
Es90157F
Note on D8: Dimensions EN (ASME/AWWA ...) ˜ Qmax for the velocity 10m/s. A size NPS by flanges AWWA (28“ and more) is indicative only. Accurate diameters aren’t defined in the standard.
Note on D9: For wafer design, specify a type of counter flanges of your pipe (ie between which flanges the sensor will be inserted)!
Note on D15: Standard combinations of nominal pressures and dimensions: PN 25 ~ DN 6 to 10 (1/8" to 1/4") PN 40 ~ DN 15 to 50 (1/2"to 2") PN 16 ~ DN 65 to 200 (2 1/2"to 8") PN 10 ~ DN 250 to 700 (10" to 28") PN 6 ~ DN 800 to 1200 (32"to 48")
Note on D17: The maximum allowable pressure (PS max) depends on: Flange type and material, max. permissible temperature (TS max) and nominal pressure (PN).
Note on D23: This data has no predefined values. Provide the most accurate description possible
Note on D27: See sheet Pulse numbers
Note on C29: Displayed volume units
Note on C30: Displayed flow units
Note on D34: Q = Q₃
Note on D37: 4 = standard
Note on D38: Baud (Bd) is a unit of modulation rate indicating the number of changes in the state of the transmission medium per 1 second.
Note on C39: Accessories: USB / RS-485 converter, connecting cables, User Manual.
Note on D39: Through the FLOSET program, the flow meter can be connected to a computer (PC, laptop, tablet) on which Windows 7 and higher (or Linux, iOS) with JAVA 8u40 and higher is installed.
Note on C40: The calibration and the official verification is icluded.
Note on D49: Product IDO (Bill of Material No. incl. execution) is entered here by an authorized ELIS employee, if the product is in the Helios database.
Note on C56: To type in this field, double-click in the field (double click - the cursor appears ...). (New Line = Alt + Enter)

## Data (hidden)
Sensor design | Code |  |  |  | Sloupec pro zápis návrhu úprav
Flange - cold and other application | 1
Wafer - cold and other application (up to DN200/8") | 2
Flange - heat meters | 3
Wafer - heat meters (up to DN200/8") | 4
Non-standard | x
Flowmeter version and equipment |  | Přepočet m³ na US Gal
ECONOMIC compact | 0 | US gal/m³: | 264.1721
ECONOMIC remote | 1
Dimension of sensor |  | m³/h | US gal/h | US GPM (US gal/min)
DN6 ~ 1 m³/h (NPS 1/4" ~ 4.4 US GPM) | 01 | 1 | 264.1721 | 4.4029
DN8 ~ 1.8 m³/h (NPS 3/10" ~ 7.9 US GPM) | 02 | 1.8 | 475.5097 | 7.9252
DN10 ~ 2.8 m³/h (NPS 3/8" ~ 12.3 US GPM) | 03 | 2.8 | 739.6817 | 12.328
DN15 ~ 7.87 m³/h (NPS 1/2" ~ 34.6 US GPM) | 04 | 7.87 | 2079.0341 | 34.6506
DN20 ~ 12 m³/h (NPS 3/4" ~ 52.8 US GPM | 05 | 12 | 3170.0646 | 52.8344
DN25 ~ 20 m³/h (NPS 1" ~ 88.1 US GPM) | 06 | 20 | 5283.441 | 88.0574
DN32 ~ 31.25m³/h (NPS 1¼" ~ 137.6 US GPM) | 07 | 31.25 | 8255.3766 | 137.5896
DN40 ~ 50 m³/h (NPS 1½" ~ 220.1 US GPM) | 08 | 50 | 13208.6026 | 220.1434
DN50 ~ 79 m³/h (NPS 2" ~ 347.8 US GPM) | 09 | 79 | 20869.5921 | 347.8265
DN65 ~ 125 m³/h (NPS 2½" ~ 550.4 US GPM) | 10 | 125 | 33021.5065 | 550.3584
DN80 ~ 200 m³/h (NPS 3" ~ 880.6 US GPM) | 11 | 200 | 52834.4105 | 880.5735
DN100 ~ 312.5 m³/h (NPS 4" ~ 1 376 US GPM) | 12 | 312.5 | 82553.7664 | 1375.8961
DN125 ~ 500 m³/h (NPS 5" ~ 2201 US GPM) | 13 | 500 | 132086.0262 | 2201.4338
DN150 ~ 788 m³/h (NPS 6" ~ 3 469 US GPM) | 14 | 788 | 208167.5773 | 3469.4596
DN200 ~ 1 250 m³/h (NPS 8" ~ 5 504 US GPM) | 15 | 1250 | 330215.0654 | 5503.5844
DN250 ~ 2 000 m³/h (NPS 10" ~ 8 806 US GPM) | 16 | 2000 | 528344.1047 | 8805.7351
DN300 ~ 3 125 m³/h (NPS 12" ~ 13 759 US GPM) | 17 | 3125 | 825537.6636 | 13758.9611
DN350 ~ 3 125 m³/h (NPS 14" ~ 13 759 US GPM) | 18 | 3125 | 825537.6636 | 13758.9611
DN400 ~ 5 000 m³/h (NPS 16" ~ 22 014 US GPM) | 19 | 5000 | 1320860.2618 | 22014.3377
DN450 ~ 5 000 m³/h (NPS 18" ~ 22 014 US GPM) | 20 | 5000 | 1320860.2618 | 22014.3377
DN500 ~ 7 875 m³/h (NPS 20" ~ 34 673 US GPM) | 21 | 7875 | 2080354.9123 | 34672.5819
DN600 ~ 7 875 m³/h (NPS 24" ~ 34 673 US GPM) | 22 | 7875 | 2080354.9123 | 34672.5819
DN700 ~ 12 500 m³/h (NPS 28" ~ 55 036 US GPM) | 23 | 12500 | 3302150.6545 | 55035.8442
DN800 ~ 12 500 m³/h (NPS 32" ~ 55 036 US GPM) | 24 | 12500 | 3302150.6545 | 55035.8442
Non-standard | xx
Sensor flanges
EN 1092-1 | 1
ASME (ANSI) B16.5 (max. 24") | 2
AS 4087 | 3
AWWA C207 (28" and larger) | 4
BS 4504 | 5
JIS B2210 | 6
Non-standard | x
Sensor material, coating
Sensor body and flanges made of carbon steel, polyurethane coating | 1
Sensor body: Carbon steel, Flanges: Stainless steel 1.4301, polyurethane coating | 2
Complete stainless steel 1.4301 | 3
Non-standard | x
Electrodes material
Stainless steel 1.4571 | 1
Hastelloy C-276 | 2
Platina-Rhodium PtRh20 | 3
Tantalum | 4
Titanium Gr.2 ASTM B348-03 | 5
Non-standard | x
Sensor lining
Hard rubber (+5 to +80 °C, DN25 to DN800) (+41 to +176 °F for 1" to 32") | 1
Soft rubber (-20 to +80 °C, DN25 to DN800) (-4 to +176 °F for 1" to 32") | 2
Special rubber for drinking water (+5 to +80 °C, DN40 to DN800) (+41 to 176 °F for 1½" to 32") | 3
Teflon PTFE (-20 to +150 °C, DN15 to DN800) (-4 to +302 °F for 1/2" to 32") | 4
E-CTFE (-20 to +120 °C, DN300 to DN800) (-4 to +248 °F for 12" to 32") | 5
PVDF (-20 to +120 °C, DN6 to DN10) (-4 to 248 °F for 1/8" to 1/4") | 6
Non-standard | x
Sensor protection
IP 65 | 0
IP 67 | 1
Non-standard IP 68 | 2
Grounding electrode
YES | 1
No | 0
Nominal pressure EN / ASME
PN 6 | 1
PN 10 | 2
PN 16 / Class 150 | 3
PN 25 | 4
PN 40 / Class 300 | 5
Non-standard | x
Maximum operation temperature of medium (TSmax)
50 °C / 122 °F | 1
60 °C / 140 °F | 2
70 °C / 158 °F | 3
80 °C / 176 °F | 4
90 °C / 194 °F | 5
110 °C / 230 °F | 6
120 °C / 248 °F | 7
130 °C / 266 °F | 8
150 °C / 302 °F | 9
Non-standard | x
Maximum allowe pressure (PSmax)
Standard (see sheet "PS max") | 1
Non-standard | x
Power supply
85 ÷ 305 V AC | 1
8 ÷ 25 V AC | 2
10 ÷ 36 V DC | 3
Protection of the transmitter
IP 65 | 1
Type of fluid acc to PED 2014/68/EU
PED is not requested | 0
Fluids of group 1 defined in the article 4, chapter 1. C) ii) PED | 1
Fluids of group 2 defined in the article 4, chapter 1. C) ii) PED | 2
Gases of group 1 defined in the article 4, chapter 1. C) i) PED | 3
Gases of group 2 defined in the article 4 chapter 1. C) i) PED | 4
Non-standard | x
Category acc to PED 2014/68/EU
PED is not requested | 0
SEP (PED 0) | 1
PED I | 2
PED II | 3
PED III | 4
Measured fluid
Water | 1
Non-standard | x
Description of measured liquid
Není nabízen ze seznamu
Grounding rings
Without grounding rings | 0
Stainless steel 1.4571 | 1
Length of cables
Without cables | 0
5 m (16 ft) (standardly) | 1
10 m (33 ft) | 2
15 m (50 ft) | 3
20 m (66 ft) | 4
30 m (98 ft) | 5
40 m (131 ft) | 6
50 m (164 ft) | 7
100 m (328 ft) | 8
Non-standard | x
Pulse number
NOT requested | 0
Standard (see sheet "Pulse numbers") | 1
Non-standard | x
Units of pulse number
Žádná jednotka (impulzní číslo není požadováno) | 00
l/pulse (standardly) | 01
hl/pulse | 02
milion l/pulse | 03
m³/pulse | 04
US liquid gal/pulse | 05
milion US liquid gal/pulse | 06
Imp gal/pulse | 07
bbl/pulse | 08
bblLiq/pulse | 09
bush US/pulse | 10
in³/pulse | 11
ft³/pulse | 12
yd³/pulse | 13
Non-standard | xx
Units of volume
l | 01
hl | 02
milion l | 03
m³ (standardly) | 04
US liquid gal | 05
milion US liquid gal | 06
Imp gal | 07
bbl | 08
bblLiq | 09
bush US | 10
in³ | 11
ft³ | 12
yd³ | 13
Non-standard units | xx
Units of flow rate
l/s | 01
l/min | 02
l/h | 03
milion l/day | 04
m³/s | 05
m³/min | 06
m³/h (standardly) | 07
m³/day | 08
ft³/s | 09
ft³/min | 10
ft³/h | 11
ft³/day | 12
US liquid gal/s (GPS) | 13
US liquid gal/min (GPM) | 14
US liquid gal/hr (GPH) | 15
US liquid gal/day (GPD) | 16
milion US liquid gal/day | 17
Imp gal/s | 18
Imp gal/min | 19
Imp gal/h | 20
Imp gal/day | 21
bbl/s | 22
bbl/min | 23
bbl/h | 24
bbl/den | 25
Non-standard | xx
Numbers samples for avaraging
25 | 1
50 | 2
100 (standardly) | 3
150 | 4
200 | 5
250 | 6
Non-standard | x
Insensitivity of measurement
50 % Q₁ (standardly) | 1
35 % Q₁ | 2
Non-standard | x
Flow rate 100% (l/s)
Standard Q₃/qp | 1
Non-standard | x
Output OUT
NOT requested | 00
Frequency for Q+ 0 to 10 kHz ~ 0 to Q₃ | 01
Frequency for Q- 0 to 10 kHz ~ 0 to Q₃ | 02
Frequency for¦Q¦ (standardly) 0 to 10 kHz ~ 0 to Q₃ | 03
Frequency for Q+ 0 to 1 kHz ~ 0 to Q₃ | 04
Frequency for Q- 0 to 1 kHz ~ 0 to Q₃ | 05
Frequency for¦Q¦ 0 to 1 kHz ~ 0 to Q₃ | 06
Frequency - non-standard (Q ≠ Q₃) | x1
Pulse for Q+ pulse 5 ms, gap min. 5 ms, standard pulse number | 08
Pulse for Q- pulse 5 ms, gap min. 5 ms, standard pulse number | 09
Pulse for ¦Q¦ pulse 5 ms, gap min. 5 ms, standard pulse number | 10
Pulse - non-standard pulses, gaps or pulse number | x2
Indication of limit flow Q > Qlimit (Qlimit = Q₃ - closes, if Q > Qlimit) | 12
Indication of limit flow Q¦ > Qlimit (Qlimit = Q₃ - closes, if¦Q¦ > Qlimit) | 13
Indication of stay withour defect (It is closed, if the flow meter is without any defect) | 14
Binary positive (closed) | 15
Binary negative (closed) | 16
Permanently closed | 17
Function OUT
Položka zrušena
Interface
Not required | 0
RS-485 MODBUS RTU | 1
Address
Není nabízena ze seznamu
Speed of transfer
300 Bd | 1
600 Bd | 2
1 200 Bd | 3
2 400 Bd | 4
4 800 Bd | 5
9 600 Bd (standardly) | 6
19 200 Bd | 7
38 400 Bd | 8
56 800 Bd | 9
FLOSET 4.0 incl. accessories
NOT requested | 0
Yes | 1
Calibration
Without calibration | 0
Standard calibration acc to EN ISO 4064-1, accuracy class 2 | 1
Above standard calibration acc to EN ISO 4064-1 with accuracy ±0,5 % | 2
Above standard calibration acc to EN ISO 4064-1 with accuracy ±0,2 % | 3
Standard calibration acc to EN 1434, accuracy class 2 | 4
Above standard calibration acc to EN 1434 with accuracy ±0,5 % | 5
Above standard calibration acc to EN 1434 with accuracy ±0,2 % | 6
Non - standard calibration | x
Number of pieces
Není nabízen ze seznamu
Packing
Unpacking | 0
Standard | 1
Non-standard | x
Method of delivery og goods
Personally | 1
By shipping agent on supplier's costs | 2
By shipping agent on buyer's costs | 3
Non-standard | x
Warranty
6 months | 1
12 months (standardly) | 2
18 months | 3
24 months | 4
36 months | 5
Non-standard | x
Note on A32: Poslední dimenze dle ASME
Note on A33: Standard: Nejmenší dimenze dle AWWA
Note on A97: Nenabízí se, dosazuje se automaticky.
Note on A106: Nenabízí se, dosazuje se automaticky.
Note on A120: Nenabízí se, dosazuje se automaticky.
Note on A135: Nenabízí se, dosazuje se automaticky.
Note on A200: Měření vždy OBOUSMĚRNÉ
Note on A220: Indikace místo Diagnostika. Opravit i u ostatních TAB?

## Tech (hidden)
Objednací číslo |  |  |  | Počet nezadaných hodnot: | 38 | Nestandardní parametry - souhrn:
|  |  |  | HW část: | 16
Pozice | Kód | Název | Hodnota
0 | FN50 |  |  |  |  | Nestandardní parametry
5 |  | Sensor design | 0
6 |  | Flowmeter version and equipment | 0
- | .2- | TECHNICAL PARAMETERS | -
8 |  | Dimension of sensor DN/max flow rate | 0
9 |  | Sensor flanges | 0
10 |  | Sensor material, coating | 0
11 |  | Electrodes material | 0
12 |  | Sensor lining | 0
13 |  | Sensor protection | 0
14 |  | Grounding electrodes | 0
15 |  | Nominal pressure PN | 0
16 |  | Max operating medium temperat. (TSmax) | 0
17 |  | Maximum allowable pressure (PSmax) | 0
18 |  | Power supply | 0
19 |  | Transmitter protection class
20 |  | Type of liquid by PED 2014/68/EU | 0
21 |  | Category by PED 2014/68/EU | 0
22 |  | Type of measured liquid | 0
23 |  | Description of the measured fluid | 0
24 |  | Grounding rings | 0
25 |  | Length of cable | 0
- | - | FLOW METER SETTINGS | 0
27 |  | Pulse number | 0
28 |  | Units of pulse number | 0
29 |  | Units of volume | 0
30 |  | Flow units | 0
31 |  | Number of samples for averaging | 0
32 |  | Measurement in sensitivity | 0
33 |  | Flow 100% (l/s) | 0
34 |  | Output OUT | 0
- | - | COMMUNICATION SETTINGS | -
36 |  | Address - enter from 1 to 247 | 0
38 |  | Transmission speed | 0
39 |  | FLOSET 4.0 with accessories | 0
- | - | CALIBRATION | -
41 |  | Calibration | 0
- | - | PURCHASE CONDITIONS | -
43 |  | Number of pieces | 0
44 |  | Packaging | 0
45 |  | Delivery | 0
46 |  | Warranty | 0
- | - | RELATED REGULATIONS | -
48 |  | Number of manual flowmeter

## PS max
Overview of standard maximum allowable pressures
depending on flange type and material, maximum allowable temperature and nominal pressure
Flanges acc to EN1092-1 carbon steel | [5 empty] | Flanges acc to EN1092-1 from stainless steel 1.4301
Nominal dimension of a sensor |  | Maximum allowable pressure (PS) [bar] for |  |  |  |  | Maximum allowable pressure (PS) [bar] for
| Nominal | the maximum allowable temperature (TSmax) * |  |  |  | Nominal | the maximum allowable temperature (TSmax) *
| pressure | TSmax = 80°C (linings HR, SR, SPR) | TSmax = 110°C (lin. PTFE, E-CTFE)** | TSmax = 120°C (linings E-CTFE)** | TSmax = 150°C (linings PTFE)** | pressure | TSmax = 80°C (linings HR, SR, SPR) | TSmax = 110°C (lin. PTFE, E-CTFE)** | TSmax = 120°C (linings E-CTFE)** | TSmax = 150°C (linings PTFE)**
DN15 to DN50 | PN40 | 38.3 | 36.7 | - | 35.2 | PN40 | 37.8 | 35.6 | - | 32.7
DN65 to DN200 | PN16 | 15.3 | 14.6 | - | 14 | PN16 | 15.1 | 14.2 | - | 13.1
DN250 to DN700 | PN10 | 9.5 | 9.1 | 9 | 8.8 | PN10 | 9.4 | 8.8 | 8.6 | 8.1
DN800 to DN1200 | PN6 | 5.5 | 5.4 | 5.4 | 5.2 | PN6 | 5.6 | 5.3 | 5.2 | 4.9
Flanges ASME B16.5 carbon steel | [5 empty] | Flanges acc to ASME B16.5 from stainless steel 1.4301
Nominal dimension of a sensor |  | Maximum allowable pressure (PS) [bar] for |  |  |  |  | Maximum allowable pressure (PS) [bar] for
| Nominal | the maximum allowable temperature (TSmax) * |  |  |  | Nominal | the maximum allowable temperature (TSmax) *
| pressure | TSmax = 80°C (linings HR, SR, SPR) | TSmax = 110°C (lin. PTFE, E-CTFE)** | TSmax = 120°C (linings E-CTFE)** | TSmax = 150°C (linings PTFE)** | pressure | TSmax = 80°C (linings HR, SR, SPR) | TSmax = 110°C (lin. PTFE, E-CTFE)** | TSmax = 120°C (linings E-CTFE)** | TSmax = 150°C (linings PTFE)**
NPS ½“ to 24“ | Class 150 | 15.9 | 15.9 | 15.9 | 15.8 | Class 150 | 14.1 | 13 | 12.8 | 12
Wafer design sensors with the cover from carbon steel
Nominal dimension of a sensor |  | Maximum allowed pressure PSmax [bar]
| Nominal | for maximum allowed temperature TSmax*
| pressure | TSmax = 80°C (linings HR, SR, SPR) | TSmax = 110°C (linings PTFE)** | TSmax = 150°C (linings PTFE)**
DN20 to DN50 | PN40 | 38.3 | 36.7 | 35.2
DN65 to DN200 | PN16 | 15.3 | 14.6 | 14
Notes
* For flanges of a material group 3E0 acc to EN1092-1 for example carbon steel P245GH, P265GH.
** On request only.
Values PSmax (maximum allowed pressure) and TSmax (maximum allowed temperature), which belong to a sensor, are printed on a production label.
It is possible to deliver sensors with flanges for other types of PSmax a TSmax in view on pressure-temperature properties of a flange material under discussion with the producer only.

## Pulse numbers
Overview of pulse numbers
depending on frequency and dimension
DN [mm] | NPS | Q4 |  | Pulse number [l/pulse]
|  | (v = 10 m/s) |  | for fmax = 100 Hz |  |  | for fmax = 50 Hz (standard) |  |  | for fmax = 10 Hz |  |  | for fmax = 5 Hz
|  | [m3/h] | [l/s] | pulse | f pulse | pulse | pulse | f pulse | pulse | pulse | f pulse | pulse | pulse | f pulse | pulse
|  |  |  | volume | for Q4 | width | volume | for Q4 | width | volume | for Q4 | width | volume | for Q4 | width
|  |  |  | V | f | Pmax | V | f | Pmax | V | f | Pmax | V | f | Pmax
|  |  |  | [l/pulse] | [Hz] | [ms] | [l/pulse] | [Hz] | [ms] | [l/pulse] | [Hz] | [ms] | [l/pulse] | [Hz] | [ms]
10 |  | 2.8 | 0.77778 | 0.01 | 77.7778 | 6 | 0.1 | 7.7778 | 64 | 0.1 | 7.7778 | 64 | 1 | 0.77778 | 643
15 | ½“ | 7.87 | 2.1861 | 0.1 | 21.8611 | 23 | 0.1 | 21.8611 | 23 | 1 | 2.1861 | 229 | 1 | 2.1861 | 229
20 | ¾“ | 12 | 3.3333 | 0.1 | 33.3333 | 15 | 0.1 | 33.3333 | 15 | 1 | 3.3333 | 150 | 1 | 3.3333 | 150
25 | 1“ | 20 | 5.5556 | 0.1 | 55.5556 | 9 | 0.1 | 55.5556 | 9 | 1 | 5.5556 | 90 | 1 | 5.5556 | 90
32 | 1 ¼“ | 31.25 | 8.6806 | 0.1 | 86.8056 | 6 | 1 | 8.6806 | 58 | 1 | 8.6806 | 58 | 10 | 0.86806 | 576
40 | 1 ½“ | 50 | 13.8889 | 1 | 13.8889 | 36 | 1 | 13.8889 | 36 | 10 | 1.3889 | 360 | 10 | 1.3889 | 360
50 | 2“ | 79 | 21.9444 | 1 | 21.9444 | 23 | 1 | 21.9444 | 23 | 10 | 2.1944 | 228 | 10 | 2.1944 | 228
65 | 2 ½“ | 125 | 34.7222 | 1 | 34.7222 | 14 | 1 | 34.7222 | 14 | 10 | 3.4722 | 144 | 10 | 3.4722 | 144
80 | 3“ | 200 | 55.5556 | 1 | 55.5556 | 9 | 1 | 55.5556 | 9 | 10 | 5.5556 | 90 | 10 | 5.5556 | 90
100 | 4“ | 312.5 | 86.8056 | 1 | 86.8056 | 6 | 10 | 8.6806 | 58 | 10 | 8.6806 | 58 | 100 | 0.86806 | 576
125 | 5“ | 500 | 138.8889 | 10 | 13.8889 | 36 | 10 | 13.8889 | 36 | 100 | 1.3889 | 360 | 100 | 1.3889 | 360
150 | 6“ | 788 | 218.8889 | 10 | 21.8889 | 23 | 10 | 21.8889 | 23 | 100 | 2.1889 | 228 | 100 | 2.1889 | 228
200 | 8“ | 1250 | 347.2222 | 10 | 34.7222 | 14 | 10 | 34.7222 | 14 | 100 | 3.4722 | 144 | 100 | 3.4722 | 144
250 | 10“ | 2000 | 555.5556 | 10 | 55.5556 | 9 | 10 | 55.5556 | 9 | 100 | 5.5556 | 90 | 100 | 5.5556 | 90
300 | 12“ | 3125 | 868.0556 | 10 | 86.8056 | 6 | 100 | 8.6806 | 58 | 100 | 8.6806 | 58 | 1000 | 0.86806 | 576
350 | 14“ | 3125 | 868.0556 | 10 | 86.8056 | 6 | 100 | 8.6806 | 58 | 100 | 8.6806 | 58 | 1000 | 0.86806 | 576
400 | 16“ | 5000 | 1388.8889 | 100 | 13.8889 | 36 | 100 | 13.8889 | 36 | 1000 | 1.3889 | 360 | 1000 | 1.3889 | 360
450 | 18“ | 5000 | 1388.8889 | 100 | 13.8889 | 36 | 100 | 13.8889 | 36 | 1000 | 1.3889 | 360 | 1000 | 1.3889 | 360
500 | 20“ | 7875 | 2187.5 | 100 | 21.875 | 23 | 100 | 21.875 | 23 | 1000 | 2.1875 | 229 | 1000 | 2.1875 | 229
600 | 24“ | 7875 | 2187.5 | 100 | 21.875 | 23 | 100 | 21.875 | 23 | 1000 | 2.1875 | 229 | 1000 | 2.1875 | 229
700 | - | 12500 | 3472.2222 | 100 | 34.7222 | 14 | 100 | 34.7222 | 14 | 1000 | 3.4722 | 144 | 1000 | 3.4722 | 144
800 | - | 12500 | 3472.2222 | 100 | 34.7222 | 14 | 100 | 34.7222 | 14 | 1000 | 3.4722 | 144 | 1000 | 3.4722 | 144